## Slide 1
Lecture Unit
Introduction to Healthcare Electronics
Lakshana Mohee, Celia Gómez Molina, and Dimitris Tzagkas
Ansys Academic Development Team
Notes: This lecture unit is designed to provide an introductory overview to electronics within healthcare

## Slide 2: Learning Objectives
Intended Learning Outcomes | 
Knowledge and Understanding | Understanding fundamental concepts and applications of Electronics for Healthcare
Skills and Abilities | Ability to simulate two use cases: a) Antenna effect on a human head and b) Magnetic Resonance Imaging (MRI)
Values and Attitudes | Awareness of regulations, power limits, impact of Radio Frequency (RF) on human body
Resources
  Adair ER, Black DR, “Thermoregulatory responses to RF energy absorption”. Bioelectromagnetics, 2003; 24(S6):S17-S38
  J. Mologni et al., “Comments on the FCC approval of finite element method for biomedical transmitters”, SBMO/IEEE MTT-S International Microwave and Optoelectronics Conference Proceedings. 565-569. 2011. 10.1109/IMOC.2011.6169268.
  Specific Absorption Rate (SAR) for Cellular Telephones | Federal Communications Commission (fcc.gov)
  Software: Ansys Electronics Desktop v24R1 (please note that all project files are available in the downloaded folder)
Built-in Assessment
  Simulation instructions
  Group discussions
  Critical thinking pointers
Further reading/information
  Ansys Academic
  Ansys Innovation Courses
  Ansys Education Resources
Notes: The learning objectives for this lecture unit are mentioned here. The main goal is to provide an understanding of the fundamental concepts and applications on Electronics for Healthcare. We will introduce the concept of Specific Absorption Rate which will be helpful to carry out our simulations for two use cases: the effect of an antenna on a human head and magnetic resonance imaging. The lecture unit also contains elements of regulations, power limits and prompts the reader to think about the safety of radio frequency on the human body.

## Slide 3: Introduction
This lecture unit introduces Electronics in Healthcare, covering applications and safety measures related to Specific Absorption Rate (SAR). Additionally, it showcases two case studies demonstrating the simulation of SAR using Ansys HFSS. Specifically, it analyzes the impact of an antenna placed near the head and an MRI system.
Pre-requisites:
Having the basics of Electromagnetics and beginner knowledge of Ansys HFSS.
This resource should be considered after completing Intro to Ansys HFSS and the Microstrip Patch Antenna Array courses.

## Slide 4: Outline
Electronics for Healthcare
Specific Absorption Rate (SAR)
Electronics with Ansys
Two Case Studies:
  Effect of Cell Phone Antenna on Human Head
  MRI safety
Closing Remarks
Notes: The outline of the presentation is mentioned here

## Slide 5: Electronics for Healthcare
Notes: This first section will discuss the revolution of electronics in healthcare

## Slide 6: Digital Transformation of Healthcare
Enhances quality, accessibility and efficiency of services
Remote consultations break down barriers
Digital records streamline workflows and resource allocation
Notes: The healthcare industry has encountered a digital transformation throughout the years. This refers to the strategic adoption of digital technologies and processes to improve various aspects of the healthcare industry such as enhancing the quality of services. Moreover, remote consultations break down barriers such as waiting time and geographical location, making healthcare more accessible. Another example is the use of digital records that make the workflow between doctor and patient more efficient.

## Slide 7: Some Electronics Examples in Healthcare
Wearables
Glucose Monitors
Robotics
ECG devices
Deep Brain Stimulation
Hearing Aids
Notes: Electronics have become an important part of established and commonly used medical innovations around the world. From patient monitors to diagnostic imaging systems and infusion devices, the use of electronics is widely spread. In this slide, we will try to look at six of the popular medical innovations that depend on the latest electronic and electrical components to successfully improve patient lives.
Wearables: The most widely known modern innovation in which electronics are playing a fundamental role is wearables. This includes glasses, smart watches, rings, and even clothing items with promising patient monitoring capabilities. The development of novel biosensors for measuring things such as heart rate, body temperature, and sleeping habits is the focus of wearables. Other than the basic functionality that they offer, there are also many other factors for manufacturers to consider when it comes to choosing the electronic components for a wearable. The innovations are said to be comfortable, discreet, energy-efficient, robust, lightweight, etc. Other than allowing people to measure their vitals on the go, wearables drive the uptake of telemedicine, which has improved social distancing protocols during the pandemic.
Glucose monitors: These continually monitor your blood glucose (blood sugar) in real time. These devices are attached to your body and provide updates throughout the day. They have become more accurate over the years and are now considered a viable treatment option for people with diabetes.
ECG devices: These monitors are essential for tracking heart health.
Brain Stimulation Devices: Brain stimulation is used for treating mental health problems and is a method that dates as far back as the 1700s. The devices used for delivering a low-level pulse of electricity to the brain and stimulating neural activity become more advanced as developers strive to improve the safety and performance of technologies. A lightweight and better-powered headset has been developed to administer direct current stimulation for redressing imbalances in the brain thought to cause mental health conditions such as depression. Many products claim to improve the cognitive performance, sleep quality, and general mood of the user using concepts such as neurostimulation. These products are clinically proven medical devices that have a CE mark and accreditation.
Hearing Aids: Hearing aids are yet another example of a medical technology that continues to be developed and is also subject to consistent innovation. Microelectronics play an important role in the rise of discreet and near-invisible earpieces, scaling down further than ever before. Some other capabilities being introduced are reduced power consumption, masking features, connecting to a smartphone, and even adding features such as answering phone calls and streamlining audio from TVs. Regardless of whether it is with an onboard deep neural network or the addition of a microphone, hearing aid products continue to push the boundaries of what is likely to occur by improving the user’s hearing experience.
Robotics: Robotics play many different roles in the healthcare sector, and electronics are required for each one of the functions. Robotic surgery allows many operations to be carried out with precision, control, and flexibility. This gives rise to a wealth of minimally invasive procedures that reduce recovery time and also mitigate many of the risks associated with open surgeries.
AI-powered robots are also making their way into hospitals to help doctors and nurses with everyday activities. However, the area where more advanced electronics are likely to play an important role is in prosthetics and, further into the future, in exoskeletons. Passive prosthetic body parts such as hands and feet with limited real-life functionality are beginning to give way to myoelectric and bionic limbs that are driven by external and electric power sources, responding to the user’s muscle movement and carrying out many tasks. The modernization of these assistive technologies is also more than likely to involve mind-controlled prosthetics and exoskeletons, many of which have shown great scope.

## Slide 8: Scalability of Electronics
Implantable Hearing Device
Magnetic Resonance Imaging (MRI)
Notes: As we have seen in the previous slide, electronics are present across a wide spectrum of applications. The scalability of electronics also makes it very advantages when designing medical devices. Microelectronics, such as an implantable hearing device, come with the ability to scale down the components inside the devices, which further proves beneficial because manufacturers strive to meet the demands of technology users. This generally requires them to be comfortable and simple while carrying out various functions. On the other side of the spectrum, we have the magnetic resonance imaging system which will be discussed in more detail later on in the lecture.

## Slide 9: Specific Absorption Rate (SAR)
Notes: This section focuses on the specific absorption rate also know as SAR

## Slide 10: Impact of Electromagnetic Waves
Application | 
Direct | Indirect
Pacemakers | Imaging (Xray, MRI, ultrasound)
Glucose monitors | Radiotherapy treatment
Brain neuron implants | Wearables (Heartbeat, blood pressure, oximeter, bone conduction headphones)
Telemetry and remote monitoring | Analysis of blood/tissue
How can we measure the interaction between the human body and the electromagnetic waves?
Specific Absorption Rate (SAR)
Notes: There are nowadays several medical devices and therapies that use electromagnetic waves and interact directly or indirectly with the human body. We have listed these accordingly on this slide. Some direct applications include pacemakers while indirect exposure includes medical imaging techniques.
How can we measure this interaction? The standard parameter that quantifies the interaction between the electromagnetic waves and the body is called SAR: Specific Absorption Rate.

## Slide 11: What is SAR and Why is it Important?
The Specific Absorption Rate (SAR) is a measure that quantifies the rate at which radiofrequency energy is absorbed by the human body when exposed to electromagnetic fields.
It is typically expressed in watts per kilogram (W/kg).
[1] Adair ER, Black DR. Thermoregulatory responses to RF energy absorption”, Bioelectromagnetics 2003; 24(S6):S17-S38
SAR is used to study the effects of exposure to electromagnetic fields, such as those generated by wireless communication devices including mobile phones, tablets, and similar devices. SAR limits are set to ensure that exposure to these fields does not exceed levels considered safe for human health.
Safety consideration
Compliance with Regulations
Optimizing Performance
Group discussion
Explore "Regulations" to understand the regulatory pathway for in silico medical devices

## Slide 12: Computing the SAR Parameter
SAR parameter formula:
where σ and ρ are electric conductivity and density of the biological tissue, respectively, |E| is the maximal value of the electric field induced in the human body.
How SAR can be considered in
real-world designs?
By carefully managing SAR levels, designers can create devices that deliver reliable functionality while minimizing potential health risks associated with RF exposure.
Through simulation
Notes: The use of simulation for SAR analysis in medical devices provides a safer, more cost-effective, and efficient means of ensuring that these devices operate within safe exposure limits, thereby protecting patient health and aiding in the development and regulatory approval process.

## Slide 13: FEM for Biomedical Applications
[2] J. Mologni et al., “Comments on the FCC approval of finite element method for biomedical transmitters”. SBMO/IEEE MTT-S International Microwave and Optoelectronics Conference Proceedings. 565-569. 2011. 10.1109/IMOC.2011.6169268.
Finite Element Method (FEM) has been ruled as a valid technique to simulate transmitters that are placed inside, on the surface, or near the human body by the Federal Communications Commission (FCC). Thus, FEM simulations can be employed on biomedical engineering, as in the following examples.
Case Study 1
Effect of an Antenna on Human Head
Case Study 2
MRI Safety

## Slide 14: Electronics with Ansys

## Slide 15: Ansys Electronics Desktop (AEDT)
The Ansys Electronics Desktop is a platform that enables true electronics system design. AEDT provides access to the Ansys industry-standard electromagnetics simulation solutions such as Ansys HFSS, Ansys Maxwell, Ansys Q3D Extractor, Ansys SIwave, and Ansys Icepak.
Ansys HFSS
High-frequency designs
Ansys Maxwell
Low-frequency designs
Circuit Design
Circuits analysis
Icepak
Thermal analysis
Notes: Ansys Electronics Desktop is the platform for simulating electronics devices. For these two case studies we will use Ansys HFSS.

## Slide 16: Ansys HFSS - Interface
Design window
Ribbon Tab
Project Manager window
Properties window
Message Manager window
Progress window

## Slide 17: Ansys HFSS – General Workflow
Geometry and materials
Boundary Conditions
Excitations
Solution setup and Analysis
Post-processing
Validation
Project Name
Design Name
General Simulation Workflow:
Generate 3D structure
Apply boundaries and excitations
Solution Setup
Solve
Post process
Notes: The general workflow is as follows:
create the geometry and assign the materials,
apply boundary conditions,
define the excitations,
create the solution setup
and analyze.
The final steps are post-processing for visualizing the results and the validation from the expert/engineer. This workflow is also represented in the design tree.

## Slide 18: Case Study 1
Effect of an Antenna on a Human Head

## Slide 19: SAR Level Regulation
The Federal Communication Commission (FCC) limit for public exposure from cellular phones is a SAR level of 1.6 watts per kilogram (1.6 W/kg).
Covers all frequencies and band used by the cell phone (2-5G, Wi-Fi, Bluetooth) e.g.
  2G Networks  ~900MHz , ~1800MHz
  3G Networks  ~800MHz, ~2100MHz
  4G Networks  ~800MHz, ~1800MHz, ~2100MHz, ~2600MHz
  5G Networks  ~600MHz, ~2500MHz, ~3500MHz, ~24GHz, ~28GHz, ~39GHz
Considering that the cell phone will be very close to the human body and skin
  Each human tissue has different characteristics
  Different phone mode may operate at different power level (standby, calling in good reception, bad reception area, etc)
  Different antennas have different directivities
[3] Specific Absorption Rate (SAR) for Cellular Telephones | Federal Communications Commission (fcc.gov)
Notes: Let’s take as an example an antenna that is used in a cell phone. Since it transmits electromagnetic energy near humans, a SAR (Specific Absorption Rate) analysis is required. In addition, SAR is regulated by different regulation bodies, with the example of FCC for public exposure from cellular phones at a level of 1.6 watts per kilogram. This covers all the different frequencies that a cell phone may use from 2G signal to the latest 5G signal networks, including wifi and Bluetooth networks.
SAR is dependent on the tissue characteristics, but also on the mode the device/phone is operating under. Thus, a variety of scenarios need to be considered and analyzed.

## Slide 20: Simulating SAR in HFSS
Antenna model
B
C
D
E
A
In AEDT View Tab:
ACT Extensions  HFSS Antenna Toolkit  Rectangular Inset Fed
In AEDT View Tab:
Components Library Human Body Exteriors
Body part model
Notes: We can create an initial setup for SAR analysis using existing support in HFSS, by combining the antenna toolkit and the components library.
With the Antenna toolkit, we can create antennas of different designs, configuring the main parameters required. Following the steps A-E as shown on the left, we can create a patch antenna tuned in the preferred center frequency by defining the center frequency field (B) and the substrate height (C) fields.
For the human head model, we can navigate to the Components library and select the human head – male or female.
These models do not include in detail the internal organs and tissues such as brain, eyes, etc. An average tissue characteristic is allocated.

## Slide 21: Arrange the Simulation Geometry Configuration
Generating an antenna from HFSS Antenna Toolkit provides analysis setup too
Arrange antenna and human body model for position of interest analysis
Change Solution type to Modal from HFSS tab and Solution Type option
Notes: Having the 2 main components needed for our case, we need to place them in a way that represents a realistic case. We can use the tools available to move and place the antenna OR to create multiple instances if that is required too. In our case, we will concentrate on the single antenna scenario. When the antenna is placed to the desired location, we can analyze the setup with the pre-defined analysis that was set by default when the antenna toolkit created the antenna and all the parameters.
For better power excitation, we should use the Modal solution type, and this can be changed from the Solution Type options, as shown in the following slide.

## Slide 22: Set up the Solution Type and Excitation
Select Modal option
Delete existing excitation
Create new Excitation Lumped Port at the antenna inset
Notes: So, when we open the Solution Type options, we can select the modal option and click OK. Then we need to update the excitation.
First, we delete the previously set by default excitation and we create a new one on the antenna port. In our case it is the inset port. By clicking on that face/surface, then with the right mouse click we go to Assign Excitation and we need to assign a port, as Lumped Port – as shown in the screenshot.

## Slide 23: Add Lumped Port and Define Integration Line
Give a name if preferred and then click next
Define new integration line
Define the line on the surface/port as shown in the picture
Notes: On the first window that appears, we can provide a port name and with next we are navigated to the window to define an integration line. Where it is set up as “none”, from the drop down, we select the “New Line” option and we define it. The integration line defines the mode of the excitation and we set it up as an arrow on the surface going towards the antenna side and not the ground side.

## Slide 24: Set up the power level and Analyze
Edit Sources at the excitation the setup to define excitation power
Excitation of the antenna should be set to the right level. Some nominal values are in the range of 0-20 dBm Depending on your system you can choose the power (here 1mW) up to the appropriate representative level
The excitation power level will depend on the antenna design, the cell phone orientation, the packaging/casing material, other components in the system, use cases, etc
Notes: Last is the excitation power for the antenna. The power varies from design to design, from manufacturer to manufacturer, and from test to test. A spectrum of nominal values is considered between 0 and 20 dBm for the cell phone antenna, where this value goes higher for base stations. Indicatively, in this example the power is set to 2 mWatt. Then we need to click Apply and OK.

## Slide 25: Analyzing and Plotting SAR
Analyze the setup
Field Overlays has already the option for SAR plot [Local_SAR or Average_SAR ]
We can configure also properties of different tissues or organs
Notes: The next step is the Analysis step. With the right mouse click on the Analysis OR with right click on the analysis Setup we select “Analyze all” or “Analyze” respectively. This analysis will need some time in the range of 40-70 minutes depending on the capabilities of the computer. When the analysis is complete, HFSS has available 2 functions to estimate SAR, and these are the Local_SAR and the Average SAR.

## Slide 26: SAR Plots
Local SAR plot
Average SAR plot
Plots in logarithmic scale
Notes: In this example, the maximum SAR value estimated to be a bit above the 1.6 W/kg in the case of Local SAR. For the estimation of Average SAR, the area that reaches above the limit is only a small point. The plots are done in a logarithmic scale.
This is an analysis done with an antenna that is not in a phone, no packaging material is used, it is transmitting towards the human head and not towards a base station with a power level of 2 mW. This is a good starting point for setting up a more complex scenario, with additional cell phone components, multiple antennas, different distances from the body area, different orientation of the antenna, etc.

## Slide 27: Case Study 2
MRI Safety

## Slide 28: Principle of Magnetic Resonance Imaging (MRI) systems
Fundamentals:
  High constant magnetic field – protons will align to it
  Apply a radiofrequency pulse
    Pulse ON: protons will align to it
    Pulse OFF: protons then generate electromagnetic energy to realign with the constant magnetic field
  Depending on how quick protons release energy – differentiate the tissues
Case Study:
  Simulate body human head inside a 63.8 MHz MRI system
  Two design inside ‘Case2_MRI_project’
  Detailed instructions to analyze and obtain results
Search for more information on "How the MRI works”
Notes: Magnetic Resonance Imaging (MRI) is a medical imaging technique used to visualize detailed internal structures. The fundamental principles of MRI involve the interaction between magnetic fields and nuclear magnetic moments. Here's a summary of the key principles. MRI works by aligning hydrogen nuclei in the body using a strong magnetic field and then disturbing this alignment with radiofrequency pulses; the emitted signals during relaxation are spatially encoded using magnetic field gradients to produce detailed images.
In this case study we will simulate a human head inside an MRI system working at 63.8 MHz to measure the SAR. We will use the provided Case2_MRI_project and we will detail the instruction to obtain the simulated results.

## Slide 29: Ansys Electronics Desktop (AEDT)
Open the project ‘Case2_MRI_project’

## Slide 30: MRI system
Boundary conditions
Freq. setup
Variables
Design
Shielding metallic structure
RF coils
Excitations
3D components and object
Notes: There are two designs (MRI_system and MRI_and_head) inside the project. The first one will be to analyze the MRI system itself and in the second one we include the head of the human inside the system. As can be observed, the model and materials are already defined. Also, the boundary conditions, excitations and frequency solution and sweep are already set up.
The MRI model is built by RF coils inside the shielding metallic structure.

## Slide 31: Analyze and Visualize the Electric Field
Right click on Setup1 > Analyze
Click on Global:YZ plane and then, right-click on Field Overlays > Plot Fields > E > Mag_E
1
2
Notes: Following the steps highlighted here, the magnitude of the electric field is plotted.

## Slide 32: Visualize the Electric Field
Click on Done
By double clicking on the color scale, you can change the limits

## Slide 33: Visualize the Electric Field
Right click on Mag_E > Animate, to visualize the Electric field
Notes: And can be animated as shown.

## Slide 34: MRI with head model
Repeating the same steps outlined previously to analyze the design of the MRI together with the head model
And plot the magnitude of the electric field
Why is the intensity of the electric field very high near the coils but weak outside the cylindrical shield structure that surrounds the RF coils?
Do we expect the electric field to be weaker inside the human body?
Group discussion

## Slide 35: SAR plot
Select the head and right click on Plot Fields > Other > Average_SAR
Then, click on Done
Notes: Following these steps, we can plot the Specific Absorption Rate

## Slide 36: SAR results
What are the head SAR regulations for an MRI system?
Does this design satisfy the regulation? Will it be safe for the patient?
Group discussion
These regulations aim to prevent excessive heating and ensure patient safety during MRI procedures.

## Slide 37: Closing Remarks

## Slide 38: Highlights
It is essential to consider regulatory bodies when designing medical devices.
1
2
3
SAR is an important parameter to measure how the electromagnetic radiation impacts human tissues.
In silico testing promotes patient safety and optimizes device performance.

## Slide 39: Dive into…
Learn more about Ansys HFSS tool with
About Healthcare in Ansys visit
Healthcare Engineering Simulation
SAR simulation on MRI system with ANSYS HFSS
Computational-based MRI Thermal Injury Risk Assessment for Patients Implanted With Passive Spinal Devices
Ansys Innovation Courses | Electronics

## Slide 40: Ansys Education Resources Feedback Survey
Here at Ansys, we rely on your feedback to ensure the educational content we create is up-to-date and fits your teaching needs.
Please click the link below to fill out a short survey (~7 minutes) to help us continue to support academics around the world utilizing Ansys tools in the classroom.
Feedback Survey Link
Notes: Feedback Survey link for copy-paste method  https://ansys.typeform.com/to/jcattJI5